## Tabelle1
Level measurement uncertainty calculation, assumption: DUT does not change output VSWR below -40 dBm | Data | Intermediate results | Standard Uncertainty in dB (1 s).
Uncertainty due to Powermeter/NRV-Z4 to 2 GHz
at Level in dBm | -40 | 1e-07
Noise in pW (with filter No. 3, 2s ) | 320 | 3.2e-10
Standard Uncertainty |  |  | 0.0069543
Zero Error in pW (rectangular distribution) | 50 | 0.002172
Standard Uncertainty |  |  | 0.001254
Calibration Uncertainty (2s) in dB | 0.06
Standard Uncertainty |  |  | 0.03
Linearity Error in % Power (2s) | 0.7
Standard Uncertainty |  |  | 0.015227
Temp. Effect 18° to 28° in % Power (1s) | 0.3
Standard Uncertainty |  |  | 0.013048
Powermeter Acc. In % Power (rectangular distribution) | 0.3
Standard Uncertainty. |  |  | 0.0075287
Sum Uncertainty Powermeter |  |  | 0.037533
Uncertainties due to mismatch:
Mismatch at D.U.T output-Switch
VSWR of D.U.T | 1.5
S11 of Switch | 1.1
Standard Uncertainty. |  |  | 0.05856
Mismatch at powersensor-switch
VSWR Sensor | 1.1
S22 of Switch | 1.1
Standard Uncertainty. |  |  | 0.013943
Uncertainty due to Mismatch D.U.T Output-powersensor
VSWR DUT | 1.5
VSWR SENSOR | 1.1 |  | 0.05285
Standard Uncertainty.
Sum Standard Uncertainty due to mismatch (measurement with NRV-Z4). |  |  | 0.080105
Uncertainty S21 Measurement of path d.u.t to sensor in dB(rectangular distribution) | 0.05
Standard Uncertainty. |  |  | 0.028868
Linearity Error of FSP at digital BW
Specification in dB (Special calibration) | 0.2
Standard Uncertainty. |  |  | 0.07
Uncertainty due to reproducability of Switching
Level in dB (rectangular distribution) | 0.03
Standard Uncertainty |  |  | 0.017321
Uncertainty due to S/N at -110 dBm (without Low Noise Amp)
Measured standard uncertainty in dB with 3 s meas. time | 0.06 |  | 0.06
Combined Standard Uncertainty in dB |  |  | 0.13213
Expanded Combined Standard Uncertainty in dB (95% Confidence) |  |  | 0.25898